# IEEE 829 Test Plan: Ischemia Predictor

TP-IP-001 · Version 1.1 (results recorded) · Ischemia Detection Team · Oct 31, 2026 · @Bindushree

## 1 Test Plan Identifier

- Test Plan ID: TP-IP-001

- Project: Ischemia Predictor (ML-based heart-disease / ischemia-proxy decision-support system)

- Version: 1.1 (results of automated execution recorded in Sections 18 and 19)

- Related SRS: Software Requirements Specification for Ischemia Predictor, v1.1 (requirement IDs FR-xx, NFR-xx, SEC-xx are used throughout)

## 2 References

- Software Requirements Specification, Ischemia Predictor v1.1

- Software Architecture and Design Specification, Ischemia Predictor v1.1

- Dataset heart.csv (SRS Appendix C, TBD-1 closed)

- IEEE Std 829-2008, Software and System Test Documentation

- OWASP Application Security Verification Standard (ASVS) 4.0

## 3 Introduction

This plan defines how the Ischemia Predictor is tested. The system takes patient clinical data, preprocesses it, and obtains a prediction ("Ischemia Likely" or "Ischemia Unlikely") from ML, DL or QML microservices. The goal is to verify that every SRS requirement is met: correct functional behaviour, measurable performance, reliable and repeatable results, and strong protection of patient data. Every test case in Section 18 traces back to an SRS requirement. In version 1.1, 28 automated pytest checks have been run; the remaining cases need manual or tool-based execution (Section 18.1).

## 4 Test Items

- Web client (login, entry form, result, history and admin pages)

- Input validation module

- Web server REST API (authentication, RBAC, orchestration, logging)

- Data preprocessing pipeline

- ML, DL and QML prediction microservices

- Result display and explanation module

- Prediction history and audit log database

- Model registry and model management functions

## 5 Software Risk Issues

- Incorrect predictions (false negatives are the most serious clinically)

- Mismatch between training and inference preprocessing

- Microservice failure or timeouts

- Data leakage of patient information

- Slow responses under concurrent load

- Browser compatibility issues

- The label is a proxy: model accuracy on HeartDisease does not prove accuracy for ischemia

### 5.1 Security Validation

| Security area | Validation method | SRS requirement | Test cases |
| --- | --- | --- | --- |
| Authentication and password policy | Attempt weak passwords; inspect stored hashes in the database | SEC-01 | TC-17 |
| Brute-force protection | Submit 5 wrong passwords; verify lockout for 15 minutes | SEC-02 | TC-02 |
| Encryption in transit | Scan with an SSL/TLS tool; confirm HTTP is refused or redirected and TLS 1.2+ is enforced | SEC-03 | TC-18 |
| Access control (RBAC) | Call admin APIs with a Clinician token; expect HTTP 403 | SEC-04 | TC-11 |
| Audit logging | Perform login, prediction and deletion; verify log entries | SEC-05 | TC-19 |
| Injection and XSS | Submit SQL injection and script payloads in every input field | SEC-06 | TC-12 |
| Model integrity | Corrupt a model file; verify it is rejected at load | SEC-07 | TC-15 |
| Network isolation | Attempt direct access to a microservice from outside the internal network | SEC-08 | TC-20 |
| Rate limiting | Send more than 60 requests per minute from one user; expect throttling | SEC-09 | TC-21 |

Tools planned: OWASP ZAP (vulnerability scan), Postman (API abuse cases), browser developer tools (session and storage checks). pytest is used for the automated API-level checks.

## 6 Features to be Tested

- Login, logout and session timeout (FR-01 to FR-03)

- Patient data input and validation (FR-04 to FR-06)

- Data preprocessing (FR-07, FR-08)

- Prediction using ML, DL and QML models (FR-09, FR-15)

- Result display: label, probability, risk category, top features, disclaimer (FR-10 to FR-14)

- Reset/clear data (FR-16)

- Error handling (FR-17, FR-18)

- Prediction history (FR-19 to FR-21)

- Model management (FR-22 to FR-24)

- Performance, reliability, accuracy, usability, compatibility (NFR-01 to NFR-13)

- Security protection (SEC-01 to SEC-10)

## 7 Features not to be Tested

- Internal working of third-party libraries (scikit-learn, PennyLane)

- Hardware-level performance and quantum hardware behaviour (a simulator is used)

- External network infrastructure and hosting provider uptime

- Clinical correctness of the diagnosis itself (the tool is decision support only)

- Clinical validity of HeartDisease as a proxy for ischemia

## 8 Approach

- Unit testing: preprocessing functions, validators, risk-category logic. Measured coverage is 73% of the gateway, services and shared code (target at least 70%, NFR-10); training scripts are not included in that figure.

- Functional testing: black-box tests of every feature using valid, invalid and boundary values.

- Integration testing: web server with each microservice, and with the database.

- System testing: end-to-end workflow from login to stored result.

- Performance testing: load tests with 50 and 200 concurrent users (JMeter or Locust).

- Security testing: as described in Section 5.1.

- Model validation testing: accuracy and recall on the held-out set; repeatability over 100 runs.

- Compatibility testing: Chrome, Firefox, Edge and Safari (latest two versions).

Automated tests: 28 pytest checks (tests/test_api.py, tests/test_service.py, tests/test_tc_extra.py) run with python -m pytest -q. They use an isolated SQLite database, a temporary pipeline and either a mocked or a small stand-in model, so they need no running services. UI, load, TLS, network-isolation, browser and usability cases are executed manually or with tools. Results are compared with expected results and recorded in Section 18.1.

## 9 Item Pass/Fail Criteria

A test case passes when the actual result matches the expected result stated in Section 18; otherwise it fails. The release passes when:

- 100% of High-priority test cases pass

- at least 95% of all test cases pass

- no open critical or high-severity defects remain

- NFR-01 (95% of predictions within 3 seconds at 50 users) and NFR-06 (accuracy and recall at least 85%) are met

## 10 Suspension Criteria and Resumption Requirements

Testing is suspended if any of these occurs: the web server crashes repeatedly, a model fails to load, the database is unreachable, or more than 30% of test cases in a cycle are blocked by one defect. Testing resumes after the defect is fixed and a smoke test (login, one prediction) passes.

## 11 Test Deliverables

- Test Plan document (this document)

- Test case specification (Section 18)

- Requirements traceability matrix (Section 18.2)

- Test execution report (Section 18.1 results and Section 19)

- Defect/bug reports

- Performance and security test reports (pending)

- Final test summary (Section 20)

## 12 Remaining Test Tasks

Manual UI checks (TC-08 and the on-screen parts of TC-01, TC-07, TC-09), load tests (TC-13, TC-14), TLS scan (TC-18), network-isolation test (TC-20), browser compatibility (TC-23), usability study (TC-25), an OWASP ZAP scan, and a full 15-minute idle-session check (TC-24). Measurement of NFR-02, NFR-08 and a user-guide review are also pending.

## 13 Environmental Needs

Hardware

- Test machine: 8 GB RAM or more, Intel i5 (or equivalent) or higher

- Server for load tests: 4 CPU cores and 16 GB RAM (or a cloud/virtual machine of similar size)

Software

- Python 3.10+, scikit-learn, PennyLane, Flask, Flask-SQLAlchemy, PyJWT, bcrypt, cryptography

- SQLite (local) and PostgreSQL 15+ (Compose), Docker and Docker Compose

- Browsers: Chrome, Firefox, Edge, Safari

- Tools: pytest and pytest-cov (used), Postman, OWASP ZAP, JMeter or Locust (planned)

- Test data: the public, de-identified heart.csv with a held-out 20% split (models/heldout.csv); no real patient identities

## 14 Staffing and Training Needs

Testing is done by the development team. Testers need basic knowledge of Python, REST APIs, machine-learning metrics and the security tools listed above. A short walkthrough of the SRS and the tools is required before testing begins.

## 15 Responsibilities

- Developer: implement features, fix defects, write unit tests

- ML Engineer: provide validated models; run model validation tests

- Tester: prepare data, execute test cases, log defects

- Security tester: run security validation (Section 5.1)

- Project Guide: review and approve the test process and results

## 16 Schedule

| Task | Duration |
| --- | --- |
| Test plan preparation | 2 days |
| Test case design | 3 days |
| Test environment and data setup | 2 days |
| Test execution (functional and integration) | 4 days |
| Performance and security testing | 3 days |
| Bug fixing and regression testing | 3 days |
| Final testing and test summary | 2 days |

## 17 Planning Risks and Contingencies

| Risk | Contingency and outcome |
| --- | --- |
| Dataset inconsistencies or too little validation data | Keep a backup copy of the dataset; use cross-validation if the held-out set is small. Outcome: 5-fold cross-validation was added to the training report |
| Model does not meet the 85% targets | Retrain or tune; report the achieved values. Outcome: this happened for QML (78.8% accuracy, 83.3% recall); it is reported and not served |
| QML service too slow on the simulator | Test with a reduced load for QML and record the limitation |
| Test environment downtime | Containerize the setup so it can be rebuilt quickly |
| Schedule slip | Run High-priority test cases first. Outcome: automated cases were run first; load, TLS, browser and usability cases remain |

## 18 Test Case Specification

### 18.1 Test Cases and Results

There are 25 test cases covering functional (F), non-functional (NF) and security (S) requirements. The Actual Result column names the pytest check that produced the result (all run on 31 Oct 2026, 28 checks passing). "Not run" means the case needs a manual or tool-based run that has not yet been done.

| TC ID | Type | Description and input | Expected result | SRS req | Actual result | Status |
| --- | --- | --- | --- | --- | --- | --- |
| TC-01 | F | Log in with a valid Clinician username and password | Login succeeds; entry form opens | FR-01 | test_tc01_login: HTTP 200 with token. Form display is a manual UI check | Pass (API) |
| TC-02 | S | Enter a wrong password 5 times in a row, then a correct one | Account locked for 15 minutes; correct password is refused during lockout | SEC-02 | test_tc02_lockout: fifth failure returns 423; correct password then also 423. The 15-minute expiry was not waited out | Pass |
| TC-03 | F | Submit the form with the cholesterol field empty and the age field set to "abc" | Both fields listed with valid ranges in one response within 1 second; no prediction call made | FR-05, FR-06 | test_tc03_exact_empty_cholesterol_and_bad_age: 400, both fields listed, prediction service not called. The 1-second timing was not measured | Pass |
| TC-04 | F | Submit resting blood pressure = 300 mmHg (valid range 50–250) | Value rejected with a range message; no prediction | FR-05 | test_tc04_exact_resting_bp_300: 400, message contains 50 and 250 | Pass |
| TC-05 | F | Submit a valid complete record | Exactly one label and a probability with one decimal place are shown | FR-09, FR-10, FR-11 | test_tc05_07_prediction (mocked model): 201 with one label and probability. Real ML and DL models also answered through the running web server. QML is not served (unvalidated) | Pass (ML, DL); QML blocked |
| TC-06 | F | Use inputs giving probabilities of about 20%, 50% and 80% (mocked model) | Risk shown as Low, Moderate and High respectively | FR-12 | test_tc06_risk_levels: Low, Moderate, High | Pass |
| TC-07 | F | Open any result page | Top 5 contributing features with direction, and the medical disclaimer, are shown | FR-13, FR-14, NFR-04, NFR-12 | test_tc05_07_prediction: 5 features and disclaimer in the response. On-screen display is a manual check | Pass (API) |
| TC-08 | F | Fill the form, run a prediction, click Clear | All fields and the displayed result are emptied with one click | FR-16 | Manual UI test not yet done | Not run |
| TC-09 | F | Stop the ML microservice, then submit a valid record | Message appears within 10 seconds; entered data is kept; error logged without patient identifiers | FR-17, FR-18 | test_tc09_message_text, test_tc09_service_down: 503, exact message, error body has only code, message and request ID. "Data kept" in the form is a manual check | Pass (API) |
| TC-10 | F | Run two predictions as Clinician A; log in as Clinician B and open History | A sees both records (newest first, 20 per page, with model version); B does not see them | FR-19, FR-20, FR-24 | test_tc10_history_isolation: A sees 1 record, B sees 0 | Pass |
| TC-11 | S | Call an admin API (create user) with a Clinician token | HTTP 403 and no change in data | FR-02, SEC-04 | test_tc11_rbac: 403 | Pass |
| TC-12 | S | Enter ' OR '1'='1 and <script>alert(1)</script> in every text field and API parameter | Input rejected or sanitized; no database error, no script executes | SEC-06 | test_tc12_injection_strings: login with SQL string returns 401; script in a numeric field returns 400; text in patient_ref is stored safely; database intact. Browser rendering and an OWASP ZAP scan not yet done | Pass (API) |
| TC-13 | NF | Load test: 50 concurrent users submit valid records for 10 minutes | At least 95% of predictions return within 3 seconds | NFR-01 | Load tool not yet run | Not run |
| TC-14 | NF | Load test: 200 concurrent users, compare with the 50-user result | Response time rises by no more than 25% | NFR-03 | Load tool not yet run | Not run |
| TC-15 | S | Alter one byte of the model file and restart the service | Model fails checksum verification and is not loaded; error logged | SEC-07, FR-22 | test_tampered_model_blocked: tampered file refused, 503, warning logged | Pass |
| TC-16 | NF | Submit the same record 100 times with the same model version | All 100 outputs are identical | NFR-07 | test_tc16_repeatability_100_runs: 1 distinct output over 100 calls (stand-in model; repeat once with the real models) | Pass |
| TC-17 | S | Register a user with a 6-character password; then inspect the password column | Weak password rejected; stored passwords are salted hashes | SEC-01 | test_tc17_password_policy: 400. test_tc17_hash_is_bcrypt: stored value starts with the bcrypt prefix and is not the password | Pass |
| TC-18 | S | Open the site over HTTP and scan the server with an SSL/TLS tool | HTTP refused or redirected; only TLS 1.2 or higher accepted | SEC-03 | The development server has no TLS; needs a reverse proxy to test | Not run |
| TC-19 | S | Perform login, prediction, record deletion and model activation | Each action creates an audit entry with user, action, timestamp and IP | SEC-05 | test_tc19_audit_entries: login, prediction, deletion and model_activation entries all present with the four fields | Pass |
| TC-20 | S | Call the ML microservice endpoint directly from a machine outside the internal network | Connection refused or blocked | SEC-08 | Needs the Docker Compose or Kubernetes network; not yet run | Not run |
| TC-21 | S | Send 70 requests in one minute as one user | Requests above 60 per minute are throttled (HTTP 429) | SEC-09 | test_tc21_rate_limit, test_tc21_retry_after_header: 429 with Retry-After: 60 | Pass |
| TC-22 | NF | Evaluate the active model on the held-out validation set | Accuracy at least 85% and recall at least 85% | NFR-06 | ML 90.2% accuracy and 93.1% recall; DL 87.0% and 91.2%; QML 78.8% and 83.3%. Only validated models are served (test_tc22_only_validated_models_used) | Partial: ML, DL pass; QML fails |
| TC-23 | NF | Open and run a full prediction on Chrome, Firefox, Edge and Safari | Layout and results are correct on all four | NFR-11 | Not yet run | Not run |
| TC-24 | F | Stay inactive for 15 minutes after login | Session ends and login is required again | FR-03 | test_tc24_session_expiry: expired token refused with 401; new token lifetime is about 15 minutes. A real 15-minute idle wait was not done | Pass |
| TC-25 | NF | Give 10 new clinicians a 10-minute introduction, then ask each to complete one prediction | At least 9 of 10 finish unaided in under 2 minutes | NFR-09 | Usability study not yet run | Not run |

### 18.2 Requirements Traceability Matrix

| SRS requirement | Test case(s) |
| --- | --- |
| FR-01 | TC-01 |
| FR-02 | TC-11 |
| FR-03 | TC-24 |
| FR-04 | TC-05 |
| FR-05 | TC-03, TC-04 |
| FR-06 | TC-03 |
| FR-07, FR-08 | TC-05 (preprocessing check), TC-16 |
| FR-09, FR-10, FR-11 | TC-05 |
| FR-12 | TC-06 |
| FR-13, FR-14 | TC-07 |
| FR-15 | TC-05 (run once per model: ML, DL; QML blocked while unvalidated) |
| FR-16 | TC-08 |
| FR-17, FR-18 | TC-09 |
| FR-19, FR-20, FR-24 | TC-10 |
| FR-21 | TC-19 (deletion step) |
| FR-22 | TC-15, TC-22 |
| FR-23 | TC-15 (activation step), TC-19 |
| NFR-01 | TC-13 |
| NFR-03 | TC-14 |
| NFR-04, NFR-12 | TC-07 |
| NFR-06 | TC-22 |
| NFR-07 | TC-16 |
| NFR-09 | TC-25 |
| NFR-11 | TC-23 |
| SEC-01 | TC-17 |
| SEC-02 | TC-02 |
| SEC-03 | TC-18 |
| SEC-04 | TC-11 |
| SEC-05 | TC-19 |
| SEC-06 | TC-12 |
| SEC-07 | TC-15 |
| SEC-08 | TC-20 |
| SEC-09 | TC-21 |

Not covered by a dedicated case: NFR-02, NFR-05, NFR-08, NFR-10, NFR-13, SEC-10 and the full FR-23. NFR-10 was checked by measuring coverage (73%); NFR-05 is checked by test_tc05_07_prediction and the stored record fields; the rest are verified by inspection or remain open and are reported in SRS Appendix D.

## 19 Test Execution Summary

| Measure | Result |
| --- | --- |
| Total test cases designed | 25 |
| Test cases executed | 18 |
| Passed | 17 |
| Failed or partial | 1 (TC-22: QML below 85%; ML and DL pass) |
| Not run | 7 (TC-08, TC-13, TC-14, TC-18, TC-20, TC-23, TC-25) |
| Success rate of executed cases | 94% (17 of 18) |
| Automated checks | 28 pytest checks, all passing; coverage 73% of gateway, services and shared code |

Several "Pass" results are API-level only (TC-01, TC-07, TC-09, TC-12); their on-screen parts are listed in Section 12.

## 20 Conclusion

The automated checks show that login, lockout, validation, risk categories, history isolation, role-based access, password hashing, audit logging, rate limiting, session expiry and model-integrity protection work as specified. The ML and DL models meet the 85% accuracy and recall targets (NFR-06) on the held-out set; the QML model does not (78.8% and 83.3%) and is therefore not served.

The release criteria in Section 9 are not yet fully met: NFR-06 is not met for QML, NFR-01 and NFR-03 have not been measured, and 7 of 25 cases are unexecuted. The recommended release is therefore ML and DL only, with QML documented as a known limitation, after the manual UI, load, TLS, browser and usability cases are executed. The label is a proxy for ischemia, so no claim of clinical validity is made.

## 21 Approvals

| Name | Role | Signature / Date |
| --- | --- | --- |
|  | Developer |  |
|  | Tester |  |
|  | Project Guide / Supervisor |  |